## Slide 1: Observer performance aka “ROC analysis”
Dev P. Chakraborty, Ph.D.

## Slide 2: Outline
Data collection methods
ROC method & AUC
Analyzing ROC data

## Slide 3
Data Collection

## Slide 4: I will focus on the ROC data collection method

## Slide 5: Truth vs. perceived Truth (i.e., Decision)

## Slide 6: Binary task
Truth T, actual or true disease state of case
  T = 1, actually non-diseased case
  T = 2, actually diseased case
Decision D, or perceived truth
  D = 1, case perceived as non-diseased
  D = 2, case perceived as diseased

## Slide 7: Binary task ...
Observer shown non-diseased (T = 1) and diseased (T = 2) cases, one at a time
Decision: observer classifies each case as either non-diseased (D = 1) or diseased (D = 2)

## Slide 8: Events
TN = True Negative
FN = False Negative (“miss”)
FP = False Positive (“false alarm”)
TP = True Positive (“hit”)

## Slide 9: Sensitivity and Specificity
P(X|Y) = probability of X given Y
True positive fraction (TPF)
TPF = P(D = 2|T = 2) = Sensitivity
False positive fraction (FPF)
FPF = P(D = 2|T = 1)
Specificity = P(D = 1|T = 1) = 1 – FPF

## Slide 10: Meanings
Sensitivity: how good observer is at diagnosing diseased cases
Specificity: how good observer is at diagnosing non-diseased cases

## Slide 11: Estimating TPF, FPF
N(X) means number of X
TPF = N(D = 2) / N(T = 2)
FPF = N(D = 2) / N(T = 1)

## Slide 12: Decision threshold z
Confidence level in decision varies from case to case
An “easy” case appears so obviously non-diseased or diseased that observer has high confidence in decision
In “difficult” case observer’s confidence level is lower

## Slide 13: Decision rule
If confidence level > z observer reports case as diseased
Otherwise, case is reported as non-diseased

## Slide 14: Decision rule ...
TPF and FPF can be varied by changing z
If z is increased, both TPF and FPF decrease

## Slide 15: ROC curve
Receiver Operating Characteristic
Vary decision threshold z
ROC curve is Plot of TPF vs. FPF

## Slide 16: Ratings method
Observer rates (i.e., labels) each case according to confidence that it is diseased, e.g.:
  1 = very confident non-diseased
  2 = somewhat confident non-diseased
  3 = ambiguous
  4 = somewhat confident diseased
  5 = very confident diseased

## Slide 17: Ratings method ...
5 ratings are equivalent to simultaneously running 4 binary tasks with thresholds
  z1 < z2 < z3 < z4
Yields 4 non-trivial points on ROC curve
Cumulate numbers and divide by appropriate denominator to get operating points on ROC, as in following example

## Slide 18: Example
60 non-diseased cases
50 diseased cases
Five-point rating scale
Actual clinical dataset

## Slide 19
 | Total | Counts |  |  |  | 
 |  | 1 | 2 | 3 | 4 | 5
Non-dis. | 60 | 13 | 19 | 8 | 2 | 1
Dis. | 50 | 5 | 6 | 5 | 12 | 22
 |  | Operating points |  |  |  | 
 |  | ≥5 | ≥4 | ≥3 | ≥2 | ≥1
FPF |  | 0.017 | 0.050 | 0.183 | 0.500 | 1.000
TPF |  | 0.440 | 0.680 | 0.780 | 0.900 | 1.000
1/60 = 0.017; 3/60 = 0.05; 11/60 = 0.183; 30/60 = 0.5; 60/60 = 1
22/50 = 0.44; 34/50 = 0.68; 39/50 = 0.78; 45/50 = 0.9; 50/50 = 1

## Slide 20
2+3+4+5
3+4+5
4+5
5
High z
Low z
z is decision threshold

## Slide 21: Important point!
Ratings are ordered labels, not numbers; for example, a 5-point rating scale could be:
  A = very confident non-diseased
  B = somewhat confident non-diseased
  C = ambiguous
  D = somewhat confident diseased
  E = very confident diseased
This yields identical ROC operating points as 1 through 5 scale

## Slide 22: Implication for analysis
Can only use ordering information, e.g., A < B < C < D < E
Arithmetic operations, e.g., averaging the ratings, etc., are forbidden !!

## Slide 23: AUC – area under curve
Computed from ratings
  Empirical AUC
  Analytical AUC

## Slide 24: Empirical vs. analytical AUCs
Empirical (observed) AUC
  Connect adjacent points with straight lines: this defines empirical ROC “curve”
  AUC = trapezoidal area under empirical ROC “curve”
Analytical AUC
  Fit ratings data to a model that predicts ROC curves
  AUC = area under predicted ROC curve

## Slide 25: Empirical AUC

## Slide 26: Analytical AUC

## Slide 27
K1 = N(non-diseased cases), indexed by k1
K2 = N(diseased cases), indexed by k2
The z’s are the ratings; subscripted 1 for non-diseased and 2 for diseased
Note: only ordering information is used in calculating W

## Slide 28
W = Empirical AUC
Theorem

## Slide 29: Comments
Use empirical AUC, not analytical
W = empirical AUC correspondence is used in almost all current approaches to analyzing ROC data
(Analytical AUC is not equivalent to W)

## Slide 30: Meaning of empirical AUC (or W)
Classification accuracy between non-diseased and diseased cases
0 <= AUC <= 1
AUC = 0: “worst” performance
AUC = 0.5: random performance
AUC = 1: best performance

## Slide 31: Advantage of using AUC
AUC is decision threshold independent
An important study:
  Beam et al 1996
  108 radiologist observers
  45 diseased and 34 non-diseased cases
  Binary classification task

## Slide 33: Types of ROC studies and related questions
Single observer one-modality
  One observer interprets cases in one modality
  Is single-observer AUC different from a specified value?
Multiple-observer one-modality
  Several observers interpret same cases in one modality
  Is observer-averaged AUC different from a specified value?
Multiple-observer multiple-modalities
  Several observers interpret same cases in multiple modalities
  Is observer-averaged AUC different between the modalities?

## Slide 34: AUC confidence interval
CI(AUC) = confidence interval of AUC
s(AUC) = standard deviation of AUC
95% CI(AUC)
  (AUC - 2 s(AUC), AUC + 2 s(AUC))
If specified value is outside CI, it is different from observed AUC at 5% significance level
Estimating s(AUC) requires statistics

## Slide 35: Unit normal distribution
https://dpc10ster.github.io/RJafrocRocBook/binary-task-model.html#binary-task-model-normal-distribution
Reason for factor of 2

## Slide 36: Comments on single-observer
Results apply to specific observer and cases in the study
Need to account for sampling effects:
  Observer is sampled from observer population
  Cases are sampled from case population
If sampling is accounted for results will apply to more general situations

## Slide 37: Comparing 2 modalities
Determine AUC for each modality and each observer, and associated variabilities
Statistical analysis
  Determines significance of observer-averaged difference (D AUC) in AUC between modalities
  Significance is assessed by p-value

## Slide 38: p-value
Probability D AUC or larger could occur by chance given D AUCtrue = 0
If p ≤ a (usually a = 5%) conclude AUCs are different at a-significance level

## Slide 39: Outcome of an ROC study
AUC difference is significant (if p ≤ a ) or not significant (if p > a )
If p > a cannot claim modalities have "identical" or "similar" performance

## Slide 40: Overview of software and online books
Summary of online material:
  https://dpc10ster.github.io/ai-froc-research/
Create free GitHub account to access it

## Slide 41: Getting started
Download R
  https://cran.r-project.org/
Download RStudio
  https://posit.co/downloads/

## Slide 42
Open RStudio
Create project file

## Slide 43
File -> New Project -> Existing Directory -> Browse to RJafrocDemo folder -> Create Project

## Slide 45: Load RJafroc
install.packages("RJafroc")
library(RJafroc)

## Slide 47: Getting your data into RJafroc
Input Excel file format
  https://dpc10ster.github.io/RJafrocQuickStart/quick-start-roc.html#quick-start-roc-excel
Reading the Excel file
  https://dpc10ster.github.io/RJafrocQuickStart/quick-start-roc.html#quick-start-roc-read
Analyze dataset
  https://dpc10ster.github.io/RJafrocQuickStart/quick-start-or-text.html#quick-start-or-text-analyze-dataset

## Slide 48: Read input data file
Franken1.xlsx is an Excel file containing ROC data
Read the file and create dataset object ds
ds <- DfReadDataFile("Franken1.xlsx", newExcelFileFormat = TRUE)
ds contains details about the data

## Slide 49: Analyze the dataset
ret <- StSignificanceTesting(ds, FOM = "Wilcoxon")
FOM = figure of merit or performance measure, "Wilcoxon" in this example
Results are contained in ret

## Slide 50: View the results

## Slide 51
ret$FOM

## Slide 52: Three types of randomizations
RRRC = random-reader random-case
  results apply to populations of readers and cases
FRRC = fixed-reader random-case
  results apply to specific readers and population of cases
RRFC = random-reader fixed-case
  results apply to population of readers and specific cases

## Slide 53
ret$RRRC

## Slide 54: Generating plots
PlotEmpiricalOperatingCharacteristics(ds, opChType = "ROC")
PlotEmpiricalOperatingCharacteristics(ds, trts = c(1:2), opChType = "ROC")
PlotEmpiricalOperatingCharacteristics(ds, trts = c(1:2), rdrs = 2, opChType = "ROC")

## Slide 55: My GitHub repositories
Landing page
  https://github.com/dpc10ster
Organization of repositories
  https://dpc10ster.github.io/ai-froc-research/
Focus on RJafrocQuickStart and RJafrocRocBook:
  https://dpc10ster.github.io/RJafrocQuickStart/
  https://dpc10ster.github.io/RJafrocRocBook/

## Slide 56: END