## Slide 1: Definitions
Observer(s)
  Radiologist(s)
  AI algorithm
Task
  Detecting cancers/lesions in medical cases
Performance
  A number quantifying how good the observer is at detecting lesions

## Slide 2: Tasks
Non search
  Known lesion location
    or
  Location known or implicit
Search
  Unknown lesion(s) location(s), observer needs to find lesion(s)

## Slide 3: Two types of observer performance studies
ROC (Receiver operating characteristic)
  non-search tasks
FROC (Free-response ROC)
  search tasks

## Slide 4: ROC studies
Case and image are used interchangeably
Observer shown non-diseased and diseased cases one at a time
Decisions: classify each case as non-diseased (D = 1) or diseased (D = 2)
Actually non-diseased: Truth = 1
Actually diseased: Truth = 2

## Slide 5: FOM
Decisions + Truth  FOM
FOM
  Figure of merit
  Performance measure
  Number between 0 and 1

## Slide 6: FOM examples
FOM = 1 if every case is classified correctly
  Every diseased cases is classified as diseased and every non-diseased case is classified as non-diseased
  i.e., every case is classified correctly
FOM = 0 if every case is classified incorrectly

## Slide 7: Question
What does FOM = 0.5 mean?

## Slide 8: Question?
What does FOM = 0.5 mean?
  Every diseased cases is classified at random
FOM = classification accuracy

## Slide 9: ROC studies
Truth = 0 or 1
  0 for non-diseased
  1 for diseased
Decisions + Truth  FOM
FOM
  Figure of merit
  Performance measure

## Slide 10: Truth and Decision
Truth T, true disease state status of case
  1 for non-diseased case
  2 for diseased case
Decision D, interpretation of case
  D = 1, case interpreted as non-diseased
  D = 2, case interpreted as diseased

## Slide 11: Events: TP, FP, etc.
TN = True Negative
FN = False Negative (“miss”)
FP = False Positive
TP = True Positive (“hit”)
2 x 2 Table:

## Slide 12: Sensitivity and Specificity
P(x|y) = probability of event x given y
True positive fraction (TPF)
TPF = P(D = 2|T = 2) = Sensitivity
False positive fraction (FPF)
FPF = P(D = 2|T = 1) = 1 – Specificity
Specificity = P(D = 1|T = 1)

## Slide 13: Estimating TPF, FPF
# D = 2 means number of diseased decisions
# T = 2 means number of diseased cases
TPF = (# D = 2) / (# T = 2)
FPF = (# D = 2) / (# T = 1)

## Slide 14: ROC curve
Plot of TPF vs. FPF
Since 0 <= TPF <= 1 and 0 <= FPF <= 1, the ROC curve is contained within the unit square

## Slide 15: Obtaining an ROC curve
Ratings method
Observer rates each case for confidence that it is diseased, e.g.:
  1 = very confident non-diseased
  2 = somewhat confident non-diseased
  3 = ambiguous
  4 = somewhat confident diseased
  5 = very confident diseased
Cumulate the numbers and divide by appropriate denominator to get operating points on ROC

## Slide 16
 | Total | Bin counts |  |  |  | 
 |  | 1 | 2 | 3 | 4 | 5
Non-dis | 60 | 13 | 19 | 8 | 2 | 1
Dis | 50 | 5 | 6 | 5 | 12 | 22
ROC |  | Operating points |  |  |  | 
 |  | ≥5 | ≥4 | ≥3 | ≥2 | ≥1
FPF |  | 0.017 | 0.050 | 0.183 | 0.500 | 1.000
TPF |  | 0.440 | 0.680 | 0.780 | 0.900 | 1.000
1/60 = 0.017; 3/60 = 0.05; etc
22/50 = 0.44; 34/50 = 0.68; etc
Example

## Slide 17
2+3+4+5
3+4+5
4+5
5
An operating point

## Slide 18: ROC curve
AUC = area under ROC curve
Classification accuracy between non-diseased and diseased cases
0 <= AUC <= 1
AUC = 0: “worst” performance
AUC = 0.5: random performance
AUC = 1: best performance

## Slide 19: Empirical vs. analytical AUC
Empirical AUC
  Connect adjacent points with straight lines
  Compute trapezoidal area under curve
Analytical AUC
  Fit the ratings data to a model of the ROC data
  Compute AUC under predicted ROC curve

## Slide 20: Empirical AUC

## Slide 21: Analytical AUC

## Slide 22: Comments
The modern trend is to use empirical AUC
Theorem: empirical AUC = Wilcoxon statistic W
W is probability that a diseased case is rated higher than a non-diseased case
Analytical AUC does not have such a simple interpretation , but yields other insights

## Slide 23: Moving along the ROC
Action level z
cases with confidence levels above z are called +ve and rest are called -ve
Thus z defines an operating point
As z is lowered from +∞ operating point moves up curve starting from (0,0) and approaches (1,1) as z approaches -∞

## Slide 24
2+3+4+5
3+4+5
4+5
5
High z
Low z

## Slide 25: AUC ....
AUC is action-level independent!
AUC increases
  Abn. imgs assigned high ratings (6)
  Nor. imgs assigned low ratings (1)
AUC = classification accuracy
E.g.,
  if all abn. imgs. rated ≥3 and all nor. imgs. rated <3, then AUC = 1

## Slide 26: Comparing 2 modalities
Determine AUC for each modality & each observer
Statistical analysis
  determines significance of the reader-averaged observed difference in AUC between modalities DAUCobs
  Significance ≡ p-value
  I will not bore you with details
Software: DBM-MRMC, U of Chicago, U of Iowa

## Slide 27: The p-value
Prob. DAUCobs, or larger, could occur by chance given DAUCtrue is 0
Probability is infinitesimal that random sample is exactly equal to DAUCobs; hence "or larger“
If p ≤ a (usually a = 5%) conclude
  AUCs are different at a-significance level
  Sign (+-) of DAUCobs determines which modality has higher AUC

## Slide 28: Choice of a
Convention: a = 0.05
If cost of finding a non-existent modality difference is high, choose smaller a, e.g., a = 0.01
E.g., new modality is much more expensive
Scientific honesty: a is chosen before study is conducted

## Slide 29: Outcome of a ROC study
Binary outcome
Significant (p ≤ a ) or not significant (p > a ) difference
If p > a
  Cannot claim modalities have “identical or similar” performance, or
  Study “approached” significance (e.g., p = 0.055)
If DAUCobs and clinical considerations warrant it, perform a larger study (sample size)

## Slide 30: ROC is a non-search paradigm
Binary task: case is called -ve or +ve
Applicable to tasks involving diffuse lesions, e.g., pulmonary fibrosis
Search tasks such as finding lung nodules involve not only detecting presence of nodules (binary task) but also pointing to correct locations of nodules
Not analyzing location information can lead to “faulty scoring”, e.g., giving credit when credit is not due and vice-versa

## Slide 31: An example of faulty scoring
Observer may miss a lesion and instead find a highly suspicious false lesion
The overall high rating was likely influenced by the false lesion
The ROC figure of merit rewards this type of decision, i.e., AUC increases
Bunch et al 1978 quote several other examples

## Slide 32: Right for wrong reason can have serious consequences
Missed lesion (arrow) and false positive rated 4 (cross)
Missed lesion grows
Biopsy performed at wrong location

## Slide 33: Another example
True target missed
Civilian position “taken out”
Collateral damage

## Slide 34: Consequence of not using location information
More difficult to detect a modality improvement (lower statistical power)
Missing a cancer is bad BUT ....
Missing a modality that is better at detecting cancers, which would have benefited thousands of women, is much worse
Algorithm designer may discard a “good idea”, etc.

## Slide 35: Search performance
Ubiquitous task in daily life
Cancer in the breast
Face in crowd
Contraband in luggage, etc.
To improve search performance one needs a figure of merit for search
Common methods, e.g., detection rate, do not account for action level z effect

## Slide 36: The FROC (free-response ROC) paradigm
Observer searches cases for lesions
Finds regions that are suspicious for presence of lesions
Sufficiently suspicious regions are marked and rated
Rating = degree of suspicion (e.g., 1, 2, ..., 5)
Data consists of n mark-rating pairs per case
n is a non-negative random integer, that depends on case, observer, task, imaging system, etc.

## Slide 37: Example of an FROC interpretation
One mark rated 4
In screening mammography 90% of cases have 0 marks
  Incidence of breast cancer is low (~ 0.005)

## Slide 38: True and false target localizations
If mark is close to a real target  lesion localization, LL
Other marks  non-lesion localizations, NL
Missed lesions (“FN”) are known
True negatives (“TN”) are unknown
  This is why FROC analysis has been so elusive

## Slide 39
4
4
1
NL rated 1
“FN”
Unobservable “TN”
LL rated 5
True Lesion

## Slide 40: Figure of merit for search performance
Visually summarizing search performance: FROC, AFROC and ROC curves

## Slide 41: FROC curve, 1986
LLF = # LL / # lesions
NLF = # NL / # cases
FROC curve is plot of LLF vs. NLF
Following examples are for 50 normal cases and 50 abnormal cases containing 98 lesions
Data was simulated and binned to maintain at least 5 counts in each bin

## Slide 42
 | TOT. | Bin counts |  |  |  |  | 
IMGS | 100 | 1 | 2 | 3 | 4 | 5 | 6
NL | ?? | 9 | 16 | 20 | 5 | 13 | 8
LL | 98 | 5 | 5 | 6 | 5 | 5 | 24
ROC |  | Operating points |  |  |  |  | 
 |  | ≥6 | ≥5 | ≥4 | ≥3 | ≥2 | ≥1
NLF |  | 0.08 | 0.21 | 0.26 | 0.46 | 0.62 | 0.71
LLF |  | 0.245 | 0.296 | 0.347 | 0.408 | 0.459 | 0.510
8/100 = 0.08; 21/100 = 0.21; etc
24/98 = 0.245; 29/98 = 0.296; etc
Example

## Slide 43: An FROC curve
Cannot define an action-level independent FOM
AUFCg = area under FROC curve to left of NLF = g is used by some
But g is arbitrary and there are other problems

## Slide 44: AFROC curve, 1990
Alternative FROC
FPF = # FP / # normal cases
AFROC curve is plot of LLF vs. FPF
Hybrid curve
  y-axis same as FROC
  x-axis same as ROC

## Slide 45: How does one get FPF?
ROC data can be inferred from FROC data
Cannot go the other way
Highest rating assumption
  ROC rating = rating of highest rated mark on case
FP rating = that of most suspicious NL on normal case
TP rating = that of most suspicious NL or LL, whichever is higher, on abnormal case

## Slide 46
 | Total | Bin counts |  |  |  | 
 |  | “0” | 1 | 2 | 3 | 
FP | 50 | 32 | 5 | 7 | 6 | 
LL | 98 | 48 | 10 | 6 | 34 | 
AFROC |  | Operating points |  |  |  | 
 |  | ≥3 | ≥2 | ≥1 | ≥”0” | 
NLF |  | 0.12 | 0.26 | 0.36 | 1.0 | 
LLF |  | 0.347 | 0.408 | 0.510 | 1.0 | 
6/50 = 0.12; 13/50 = 0.26; etc
34/98 = 0.347; 40/98 = 0.408; etc
Example

## Slide 47: An AFROC curve
Area q under AFROC is an action-level independent FOM
Dotted line ... later

## Slide 48: ROC curve
FPF = # FP / # normal cases
TPF = # TP / # abnormal cases

## Slide 49
 | Total | Ratings data |  |  |  | 
 |  | “0” | 1 | 2 | 3 | 
FP | 50 | 32 | 5 | 7 | 6 | 
TP | 50 | 61 | 7 | 5 | 25 | 
ROC |  | Operating points |  |  |  | 
 |  | ≥3 | ≥2 | ≥1 | ≥”0” | 
FPF |  | 0.12 | 0.26 | 0.36 | 1.0 | 
TPF |  | 0.50 | 0.60 | 0.74 | 1.0 | 
6/50 = 0.12; 13/50 = 0.26; etc
25/50 = 0.50; 30/50 = 0.60; etc
Example

## Slide 50: An ROC curve

## Slide 51: Why does ROC curve does not extend continuously to (1,1)?
Reporting an case
  assigning an ROC rating > 1 (1 is definitely normal)
  OR
  marking the case (a mark corresponds to a suspicious region)
Finite number of cases (set A) with nothing to report (~90% in mammography)
Remaining cases (set B) have something to report
As action level z is lowered, more and more cases in set B get reported until one runs out of cases

## Slide 52: Why does ROC curve not ...
cases in set A are never reported, unless
  Observer adopts guessing mechanism when there is nothing to report
  highly unlikely as mammographers do not guess
The investigator assigns the default ROC “0” rating to these cases
  in a true ROC study they would have been rated “1”
When “0s” are cumulated the non-informative (1,1) point is discontinuously reached via a jump

## Slide 53: Quantifying search performance: JAFROC

## Slide 54: JAFROC: on my website
Jackknife AFROC
FOM = area under AFROC curve (q)
q = 0% : worst performance
q = 100%: perfect performance
q < AUC; as LL is subset of TP
Dq > D(AUC); as q has a larger range
  FROC has greater modality discrimination ability
  i.e., greater statistical power

## Slide 55: JAFROC ...
Analogous to ROC software DBM-MRMC, except it applies to FROC (and ROC) data
ROC data is a special case of FROC data
Yields p-value, ...

## Slide 56: Modeling search
2-stage model (Rock, Wolfe, Kundel and Nodine, etc.)
Global-response stage (~100 ms per case)
  Observer identifies suspicious regions: noise sites, signal sites
Decision-making stage (~1 s per site)
  Observer calculates decision variable (confidence level) z
  Region is marked if z > action level z
  Rating is the bin-index of z

## Slide 57: Analogy to computer aided detection (CAD) algorithms
CAD algorithms use a 2-stage process to find lesions
Global response  "initial detections"
Decision-making  "candidate analysis“
Wasteful to give equal attention units to all parts of the case, and
Performance suffers from indiscriminate “exhaustive” search

## Slide 58
Search model parameterization
Noise site samples
Signal site samples
m
z
l
n
z
s = # of lesions
n = n1 + n2

## Slide 59: Meaning of SM parameters
l = mean number of noise sites per case
n = probability (≤1) that a lesion is a signal site (“found”)
m = ability of observer to discriminate between signal and noise sites
Performance increases for large m, small l and large n

## Slide 60: Expert
Signal site
Noise site
Lesion
Lesion
Naive

## Slide 61: Applications: Identifying the weak link
If l is large train with normal-rich cases
  Help observer/CAD reject (from stage-2 analysis) normal variants that look like lesions
If n is small train with abnormal-rich cases
  Help observer find lesions
If m is small train with abnormal-rich cases, with locations of true lesions and false lesions (for experts) shown
  Help observer recognize features that distinguish true lesions from false lesions

## Slide 62: Suggestion for CAD developers
Problem: l is usually larger than for experts
Regard suspicious regions found by experts as “pseudo-lesions”
  Train CAD to find pseudo-lesions while not finding other regions that experts rejected; this should decrease l
  Since the pseudo-lesions fooled experts, finding them may also improve n
Since normal cases are easy to acquire, CAD can be trained on a large data set
Pseudo-lesion locations obtained by asking the experts for locations of other suspicious regions they considered in their interpretation of the case

## Slide 63: Suggestion ...
Improve m by optimizing classifier performance between true lesions and the pseudo-lesions
Currently CAD developers optimize classifier performance between true lesions and false lesions found by CAD (not those found by experts)
Pseudo-lesions are more like lesions than other regions, so this type of optimization is expected to be better

## Slide 64: Estimation of SM parameters

## Slide 65: 2-step ML procedure
Likelihood function L (m, l, n, FROC data)
For values of all other parameters determine m to minimize goodness-of-fit statistic G
Effectively reduces the number of parameters to two: l and n
While (not converged) {
  Determine m to minimize G
  Vary l and n and cutoffs zi
} until L is maximized

## Slide 66: DM = digital mammography BT = breast tomosynthesis

## Slide 68: Evidence for search model
SM is not concerned with how one gets to noise and signal sites (domain of CAD)
SM makes parametric assumptions which makes math more tractable, but these could be relaxed
Direct evidence is from eye-tracking data on radiologists
Kundel and Nodine, Krupinski, ...

## Slide 69: Evidence...
SM ROC curves are proper (never cross the chance diagonal)
Inaccessible dotted portion on ROC is consistent with clinical data
Difficult to get radiologists to move sufficiently “up” the curve
Consequence of small l (~ 1 for experts, i.e., roughly 37% of cases have no noise sites)

## Slide 70: Evidence...
Binormal model (Dorfman and Alf)
  abnormal case pdf is wider than normal case pdf (b < 1)
In SM the noise and signal site pdfs are unit variance, but since TP rating could be from either distribution, the effective width of TP pdf increases
Also due to multiple noise site samples, width of FP pdf decreases
With appropriate choice of parameters SM behaves like other ROC models

## Slide 71: DEMO

## Slide 72: Recent work
Incorporating correlations into search model: intra-case, inter-modality, etc.
Fairly sophisticated FROC simulator

## Slide 73: Needed extensions of search model
Satisfaction of search (SOS)
Observer finds a lesion but does not find other lesions
Premature termination of search on lesion-rich cases
Possible explanations:
  A: n is smaller on lesion-rich cases
  B: z is higher on lesion-rich cases

## Slide 74: Needed extensions ...
Fewer NLs on abnormal cases
Possible explanations:
  A: l is smaller on abnormal cases
  B: z is higher on lesion-rich cases

## Slide 75: Needed extensions ...
Vigilance: missing rare lesions (Wolfe, Science)
Possible explanations:
  A: n decreases with decreasing abnormal case prevalence, approaching 0 as prevalence approaches 0
  B: z increases as prevalence approaches 0

## Slide 76: Avenues for research
These hypotheses can be tested
Simultaneous eye-tracking and FROC data
  Eye-tracking tells us where they looked
  FROC tells us the decisions they made
Integrated analysis of eye-tracking and FROC data could yield insight

## Slide 77: Conclusions
Understanding search is crucial to vision science and its practical applications
Widely researched field
SM comes at it from a different angle by relating it to FROC paradigm
Commonly held beliefs in non-visual science community, e.g., exhaustive search, not supported by evidence

## Slide 78: Conclusions ...
Modern tools available for FROC analysis and parameter estimation
Higher statistical power
Greater relevance to search tasks

## Slide 81: Comparing 2 modalities
Diagnostic task
Data acquisition paradigm
2 or more modalities, radiologists, cases
Ratings data
Figure of merit
One value per reader per modality
ROI
FROC
ROC
LROC
ROI: AUC
FROC: JAFROC, others
ROC: AUC, others
LROC: PCL, others
Significance testing
p-value, 95% CI
Parametric or non-parametric
DBM-MRMC, OR, BWC, others
Mammography, chest, etc.

## Slide 82: FROC figure of merit
In medical context: target = lesion
LLF = lesion localization fraction
LLF is fraction of correctly localized lesions
AFROC curve is plot of LLF vs. FPF
Area under AFROC is desired figure of merit
It is the probability that a lesion is rated higher than highest rated mark on a normal case

## Slide 83: Why observer performance?
Fourier measurements (MTF, SNR, etc.) are useful for optimizing parts of imaging chain, e.g., detector
Display, case processing, observer (human or computer) skill, the task, etc., are not accounted for
Cannot identify the “weak-link”

## Slide 84: Estimation of parameters
Expressions for ROC, FROC, AFROC curves and likelihood are available
E.g. for ROC (s is number of lesions)
MLE

## Slide 85: Likelihood function
r is a ratings bin index
P (zr < z < zr+1 | NOR) = FPF(zr) - FPF(zr+1)
P (zr < z < zr+1 | ABN) = TPF(zr) - TPF(zr+1)

## Slide 86: 2-step ML procedure
2-step method
Starting with trial values of all parameters except m
While (not converged) {
  Determine m to minimize chi-square statistic
  Vary remaining parameters to maximize likelihood
} until likelihood is maximized